Changes in TetrAMM firmware version 2.0

# The “NAQ” command does not start the acquisition
The “NAQ” command can be used to set the number of samples to be captured. It no longer starts the acquisition, the user should use ACQ and TRG commands to start the acquisition.
NAQ:? returns the number of samples to be captured.

# The “TRG” command behavior changed
With NAQ = 0, the TRG command will start to capture the samples at positive edge of trigger input signal and will stop with capture at negative edge of trigger input signal. The data transmitted include a header and a footer, see the “Trigger data format”.
With NAQ != 0, the TRG command will start to capture the samples at the positive edge of trigger input signal and will continue until the NAQ samples are captured. The data transmitted include a header and a footer, and an “ACK\r\n” is also included at the end, indicating that TRG command has finished and is in disabled state.

# The “ACQ” command behavior changed
With NAQ = 0, the ACQ command will start to capture the samples immediately and will stop with ACQ:OFF or S command.
With NAQ != 0, the ACQ command will start to capture the samples immediately and will continue until the NAQ samples are captured. The “ACK\r\n” is also included at the end.

# Data format in triggered mode changed
The format of data in the triggered mode is now always aligned to (number of ch + 1)*8 bytes. Each “line” is composed from measured data or control information and “end-of-line” signaling NaN. There are 3 distinctive “end-of-line” NaNs:
- header (0xFFF40000FFFFFFFF)
- footer (0xFFF40001FFFFFFFF)
- data (0xFFF4002FFFFFFFF)

Example of the data packet in triggered mode with ASCII formatting disabled:

fff400000000001f fff400000000001f fff400000000001f fff400000000001f fff40000ffffffff
3e3af6dfc5cd7639 be43fe414e7ea2c6 be6d96af038dc1ee be63d57689c9d2ff fff40002ffffffff
3e3ba4f00ef0d355 be43f1b8a9ffd308 be991437e6f6c24e be63be46067448da fff40002ffffffff
3e3b54690de03338 be44022ff0884d5c be9a16aee1ffdefb be63caac10945d21 fff40002ffffffff
3e3befcbd555a60d be43c8ffb8b21e02 be731f8e7ddc6172 be63bd5bab213bf0 fff40002ffffffff
3e3d06d0d498e3d6 be43e8c8abd5963d 3e94e39b34a40306 be63d077ccbff1e6 fff40002ffffffff
fff40001ffffffff fff40001ffffffff fff40001ffffffff fff40001ffffffff fff40001ffffffff

# GATE command no longer present
The TRG command has now the same functionality as GATE command, so the later was removed from the command set.

# S command during ACQ with NAQ != 0 does not emit ACK

# Buffers have adjustable size to reduce the latency